## Slide 1: A Real Time Application to Identify Chronic Alcoholics from ECG Signals
Akarsh N Kolekar (1PI13EC009) Guide: Dr. b Niranjana Krupa
Apoorv Vatsal (1PI13EC017)
Rakshith Vishwanatha (1PI13EC075)

## Slide 2: Objective
To develop a prototype to read ECG signals
Extract features from ECG signal to perform HRV analysis
Classify the person under test as an alcoholic or otherwise, in real time.

## Slide 3: Literature Survey
The paper “Effects of Alcohol on the Electrocardiogram” took electrocardiograms of 1,000 chronic alcoholic patients and analyzed them to find evidence that excessive consumption of alcohol may produce changes in the electrocardiogram.
In the “A Precise Drunk Driving Detection Using Weighted Kernel Based on Electrocardiogram” paper the authors collected ECGs from 50 volunteers. To identify characteristics of an alcoholic, features such as Pmax, Pd, means and variances of P, R, S waves and R-R intervals are extracted. The paper used SVM as a classifier and 10-fold cross validation.

## Slide 4: Literature Survey
The “Circuit Design for Front-End Electrocardiograph”, “Techniques for accurate ECG signal processing”, “Electrocardiogram (ECG) circuit for use with oscilloscopes” and “ECG Circuits, Signal Sampling and Digitalization” detailed the circuit elements that would be required in order to amplify the differential signal, filter out the different kinds of noise present in it.
In “What are Extreme Learning Machines? Filling the Gap Between Frank Rosenblatt’s Dream and John von Neumann’s Puzzle”, the author describes the architecture of ELM and compares it with SVM and LS-SVM.

## Slide 5: Physiological Effects of Alcohol
Heart rate mainly controlled by autonomic nervous system (ANS)
  2 divisions of ANS:
    Sympathetic nervous system (SNS)
    Parasympathetic nervous system (PSNS)
Alcohol consumption over time affects the SNS and PSNS
SNS activity increases heart rate (HR), which reduces heart rate variability (HRV)
PSNS activity decreases HR, which increases HRV
Alcohol suppresses PSNS activity leading to decrease in HRV

## Slide 6: HRV (cont.)
ALCOHOL
SNS
ANS
PSNS
HR
HRV
HRV:
  TDF
  FDF
  NLF
  ARX

## Slide 7: Methodology
Main verticals of the project:
  Hardware Design
      ECG signal acquisition
  Software
      Pre-processing
      Feature Extraction
      Classification
  Raspberry Pi
      Real-time, Portability

## Slide 8: Overall Flow
SVM , ELM
1) PSD

## Slide 9: Hardware
Main sensor circuit consists of:
  Instrumentation Amplifier
  High Pass Filter
  Low Pass Filter
  Notch Filter
Multiple circuit designs were simulated on Multisim.
Designed filters depending on what frequencies needed to be filtered.
Started soldering components onto circuit boards.
Alternatively, started interfacing AD8232 based sensor with Raspberry Pi. Also used an ADC MPC3008 since R-Pi doesn’t have it inbuilt.
Captured ECG samples using setup and stored it as CSV file on R-Pi.

## Slide 10: Circuit Design – 1
Low noise
High gain
High supply voltage
+ve and –ve supply

## Slide 11: Need for Notch Filter

## Slide 12: Circuit Design – 2

## Slide 13: Software – Dataset Description
Dataset obtained from Autonomic lab, NIMHANS, Bengaluru.
Total number of ECG samples = 67 of just 2 classes
Number of Alcoholic subject samples = 38
Number of Non Alcoholic subject samples = 29
ECG data obtained from test subjects at sampling frequency of 1kHz.
Each sample at NIMHANS was recorded for a duration of 5 minutes, which gave a CSV file containing 300,000 values.

## Slide 14: Software Pre-processing
ECG signal contains
  Baseline wandering
  Power-line interference
  Motion artefacts
IIR HPF (f_cut = 5Hz - 7Hz, rp = 0.2, rs = 1) to
  Remove baseline wandering
  Suppress motion artefacts
  Sufficient for only time domain and non-linear
Wavelet Transform
  Remove baseline wandering
  Remove (almost completely) all other artefacts
  Satisfies requirement of clean signal for frequency domain

## Slide 15: IIR Filtered Signal
Original Signal in Dataset
IIR Filtered Signal

## Slide 16: Wavelet Decomposed Signals

## Slide 17: Software – Feature Extraction
Types:
    Time Domain features
      Depend on R-R peak intervals
    Frequency Domain features
      Depend on Power Spectral Density
    Non Linear features
      Depend on R-R peak intervals
    ARX coefficients
      Obtained coefficients of different orders

## Slide 18: Feature Set for a Sample
[ rr_mean, rr_std, HR_mean, HR_std, rr_rms, rr_50, nrr_50, sd1, sd2, apen,
pk_freq_vlf, pk_freq_lf, pk_freq_hf, ab_pow_vlf, ab_pow_lf, ab_pow_hf, pw_ttl, rp_vlf, rp_lf, rp_hf, norm_lf, norm_hf, ratio, [ARX_coeffs] ]

## Slide 19: Visualisable Features
Poincare Plot – relates i-th RR interval with next RR interval
PSD of Wavelet Decomposed Signal

## Slide 20
Feature Extraction – ARX
  ARX Model Structure:

## Slide 21
Feature Extraction – ARX
System Identification toolbox

## Slide 22: Feature Extraction – ARX

## Slide 23: Classifiers – SVM
Decision boundary/primary hyperplane idea to separate
different classes of data
Maximize the margin between the supporting hyperplanes
Need to solving the Lagrangian Dual to obtain the weights for the primary hyperplane
Simplified SMO algorithm used to solve Lagrangian Dual
  Initialization of alpha and b not optimal
Accommodation of outliers if no regularization – choice of regularization parameter ‘C’

## Slide 24: Classifiers – SVM
Gaussian Kernel used to take samples to a new dimension for easier separation -- choice of sigma
The system was re-trained 400 times. In each loop, dataset was split randomly for classification and validation using crossvalind().
Now, for each set of randomly set classification and validation groups, all pairs of C and sigma were looped through to see which C and sigma resulted in highest accuracy.
Similarly, a loop was run 200 times to find optimal C and sigma when order 3 and order 5 ARX features were included

## Slide 25: Classifiers – SVM
C and sigma values:
  No ARX: C = 0.1, sigma = 0.3 [ 80% ]
  ARX order 3: C = 0.3, sigma = 1 [ 82% ]
  ARX order 5: C = 0.3, sigma = 1 [ 86% ]
  ARX order 5 gave better accuracy due to better concentration
To test how many TP and TF were occurring (for ARX order 5), confusion matrix was calculated after setting the C and sigma as given above:
Sensitivity was found to be = 89%
Specificity was found to be = 82%
25 | 3
5 | 23

## Slide 26: Classifiers – ELM
Feed forward neural network
Single Hidden Layer
Random assignment of Input Weights
Fixed Network Architecture
  Number of Hidden Layer Neurons fixed before training
Known for high speed of training
Also known to produce good generalisation for low number of training data

## Slide 27: ELM (cont.)
Leave-one-out validation and K-fold cross validation(k=7)
  balanced dataset of 28 alcoholic samples and 28 non-alcoholic samples (total of 56 samples).
Sigma value was varied, and the number of hidden layer neurons was varied for each sigma, and a plot of accuracy vs. number of hidden neurons was plotted.
Each training-testing iteration was done 20 times to obtain the weights and biases for which maximum accuracy was obtained.
These weights and biases were saved onto a CSV file which would then be imported into the Python classifier code, which would run on the Raspberry Pi.

## Slide 28: ELM (cont.)
Accuracy
Number of Hidden Neurons

## Slide 29: ELM (cont.)
Leave-one-out validation (with ARX order 5)
    Accuracy for Leave-one-out validation = 92.86%
    Sigma = 2
    Number of Hidden Neurons = 6
7-fold cross-validation (without ARX coefficients)
    Accuracy for 7-fold cross validation = 89.29%
    Sigma = 1.4
    Number of Hidden Neurons = 6

## Slide 30: ELM (cont.)
7-fold cross-validation (with ARX order 3)
  Accuracy for 7-fold cross validation = 92.86%
  Sigma = 2
  Number of Hidden Neurons = 7
7-fold cross-validation (with ARX order 5)
  Accuracy for 7-fold cross validation = 94.64%
  Sigma = 1.8
  Number of Hidden Neurons = 15

## Slide 31: ELM (cont.)
Confusion matrix (without ARX) was calculated as follows:
Sensitivity was found to be = 92.86%
Specificity was found to be = 85.71%
26 | 2
4 | 24

## Slide 32: Classifiers – McELM
For ELM, to find number of hidden layer neurons which gives best accuracy without over-fitting, the number of neurons was varied between a certain range.
McELM provides an intelligent way of choosing and adding neurons based on some criteria during the training phase, instead of using the brute force method to obtain an optimal solution.
We have gone through the algorithm and this is part of future work.

## Slide 33: SVM vs ELM
K-fold Cross validation:
  Without ARX:
      80% vs 89%
  With ARX order 5:
      86% vs 94%
Confusion Matrix:
vs
26 | 2
4 | 24
25 | 3
5 | 23

## Slide 34: Raspberry Pi - Python Implementation
Interfacing of sensor to R-Pi via ADC
1kHz sampling of signal implemented with timer() class to match training dataset
Code to obtain and save sampled values from sensor was written in Python.
Feature extraction codes except for the frequency domain features have been implemented in Python.
ELM and SVM Classifiers have been implemented in Python.

## Slide 35: Analog Sensor Interface to R-Pi

## Slide 36: ECG Signal Capture

## Slide 37: Summary and Future Work
  Summary:
    Classifiers can be applied to features extracted via HRV analysis
    ELM classifier has a better accuracy than SVM
    ARX features improve the accuracy of both ELM and SVM
  Future Work:
    Test the system on an alcoholic test subject
    Implement Non-Linear ARX features to see if further improvements can be made on the accuracy of the system
    Implement McELM and compare with performance of ELM
    Multi-class classifier to detect different levels of alcoholic consumption

## Slide 38: References
G.-B. Huang, “What are Extreme Learning Machines? Filling the Gap between Frank Rosenblatt's Dream and John von Neumann's Puzzle,” Cognitive Computation, vol. 7, pp. 263-278, 2015.
Andrew Ng, “Support Vector Machines”, 2011. [Online] Available: http://cs229.stanford.edu/notes/cs229-notes3.pdf Accessed: 10-Feb-2016
Andrew Ng, “The Simplified SMO Algorithm”, 2012. [Online] Available: http://cs229.stanford.edu/materials/smo.pdf Accessed: 10-Feb-2016
Branislav Vuksanovic & Mustafa Alhamdi, “AR-based Method for ECG Classification and Patient Recognition,” International Journal of Biometrics and Bioinformatics (IJBB), vol. 7 ,Issue 2, 2013
Dingfei Ge, “Cardiac arrhythmia classification using autoregressive modelling”. [Online] Available: http://biomedical-engineering-online.biomedcentral.com/articles/10.1186/1475-925X-1-5

## Slide 39: References
Pico Technoology, “Electrocardiogram (ECG) circuit for use with oscilloscopes”. [Online] Available: https://www.picotech.com/library/application-note/electrocardiogram-ecg-circuit-for-use-with-oscilloscopes
Lv Jinhua and Xu Yanyi, “Circuit Design for Front-End Electrocardiograph”, International Journal of Multimedia and Ubiquitous Engineering Vol.11, No.5 (2016), pp.345-354
Ajay Bharadwaj and Umanath Kamath, Cypress Semiconductor Corp., “Techniques for accurate ECG signal processing” [Online] Available: http://www.eetimes.com/document.asp?doc_id=1278571
NTHU, “ECG Circuits, Signal Sampling and Digitalization” [Online] Available: https://www.google.co.in/url?sa=t&rct=j&q=&esrc=s&source=web&cd=1&cad=rja&uact=8&ved=0ahUKEwins5jwg43TAhWCq48KHcxPAysQFgggMAA&url=http%3A%2F%2Flms.nthu.edu.tw%2Fsys%2Fread_attach.php%3Fid%3D692351&usg=AFQjCNFJqOwwhxZZG0FP1yQNMuNpHBV73Q&sig2=e_-aLgXrLbFMIB-XhTd2Kg

## Slide 40: References
L. Ljung,“System identification toolbox,” The Matlab user’s guide, 2012. [Online] Available: http://radio.feld.cvut.cz/matlab/pdf_doc/ident/ident.pdf
Hong He, Xiaowen Yan and Wei Wei, “Meridian ECG Information Transmission System Modeling Using NARX Neural Network” ,IEEE/ACIS 15th International Conference on Computer and Information Science (ICIS), 2016. [Online] Available:http://ieeexplore.ieee.org/document/7550775/
R. Savitha, S. Suresh, H.J. Kim, “A Meta-Cognitive Learning Algorithm for an Extreme Learning Machine Classifier” Cognitive Computation, vol. 6, pp. 253-263, 2014.

## Slide 41: Thank You